# Test and calibration reference document for the Atlas sensor suite

Sensor Reference Documents	1

# Sensor Reference Documents

The following documents contain information on the sensors and the sensor interface boards.  The information includes storage, operating, and calibration procedures - the basic care and feeding of the sensors that will need to become part of the daily use of these sensor systems.

Table 1: Sensor Reference documents

| Sensor | Probe Datasheet | Interface Datasheet | Modified Mega TWI code | Notes | Data Documents, other links | Storage Solution (SS) / calibration solutions |
| --- | --- | --- | --- | --- | --- | --- |
| DO | DO Probe | DO EZO datasheet | Do_Code | Needs calibration in the lab. Compensation necessary in the field. Procedures in EZO datasheet. |  |  |
| ORP | ORP Probe | ORP EZO Datasheet | Orp_Code | Needs calibration in the lab. No compensation req’d. Procedures in EZO datasheet. |  | SS req’d |
| pH | pH Probe | Ph EZO Datasheet | Ph_Code | Needs calibrated. Temperature compensation req’d in the field. Procedures in the EZO datasheet. | Readings Doc | SS req’d, 3xCS (pH = 4, 7, 10) |
| EC | EC Probe | EC EZO Datasheet | Ec_Code | Calibration is required - and somewhat complicated in that temperature compensation is required when calibrating. Temperature compensation req’d in the field. Procedures in the EZO datasheet. | EC Accuracy graphs |  |
| RTD (temp) | PT-1000 Temp probe | RTD EZO Datasheet | Rtd_Comp | Calibrate in boiling water. Note that water boils at 94.5 degrees C at the altitude of the Innovation Center. Procedures in the EZO datasheet. |  |  |
| Turbidity | Turbidity sensor board ckt design | Sensor design tasks |  | Sensitivity analysis |  |  |
| Color | Arduino |  |  |  |  |  |
| Light intensity | Arduino |  |  |  |  |  |

Design considerations: Water flow, mounting in pipes.

# Calibration process

## Temperature Probe

- This calibration process comes from the RTD Data Sheet Page 12.

- The innovation center is at 4942’ of elevation.  Water boils here at 95.04 degrees based on this calculation.

- Before proceeding to the lab, we tested the command to read a connected sensor with the sensor attached.

- The sensor properly connected and we read values that went up and down based on whether we were heating or cooling the probe.

- We used distilled water and boiled it, then inserted the probe. (See image below)

- For the first temperature probe, labeled A, the calibration data is:

  - Success

  - ?EXPORT,2,20

  - Success

  - 00CB21BF0000

  - Success

  - 0100283D

  - Success

  - *DONE

- Interpreting this for Temperature Probe A:

  - We should see TWO strings, representing TWENTY bytes of data.

  - What we actually see is two strings, representing 20 characters (bytes) of data.

    - Note: two one byte characters actually equals one real byte of data in memory.  They use a confusing terminology here.

  - The calibration sting in this case is: 00 CB 21 BF 00 00 01 00 28 3D

  - The Calibration import commands when using temperature sensor A should be:

    - Import, 00CB21BF0000

    - Import, 0100283D

- For the second temperature probe (probe B) the calibration data is:

  - Success

  - ?EXPORT,2,20

  - Success

  - 00CD2CBF0000

  - Success

  - 01004F20

  - Success

  - *DONE

- Interpreting this for Temperature Probe B:

  - We should see TWO strings, representing TWENTY bytes of data, which we see.

  - The calibration sting in this case is: 00 CD 2C BF 00 00 01 00 4F 20

  - The Calibration import commands when using temperature sensor B should be:

    - Import, 00CD2CBF0000

    - Import, 01004F20

- These results need to be copied to the calibration tables below.  They will also need to be put into the pre-field-trip checklist so that the sensor suite is calibrated for the sensors that are actually going into the field.

## ORP Probe

- Page 12-13 in the ORP Manual

- In the I2C mode we must continuously request readings.  This is the “R” command - send it every second and wait for a reply.

- Find the 225 mV calibration solution pouch (Bottle Not Pouch) and have it nearby.

- Rinse of the sensor first (distilled water preferred).

- Open the 225 mV calibration pouch and insert sensor into the pouch.  Take continuous measurements until the readings stabilize.  The stabilization may have slight fluctuation in values.

- Once stabilized, issue the command “cal,225”

- In order to capture the calibration values (which may be needed in the future), use the export calibration command (see page 50).  Cut and paste the output in this document.

- Orp A Calibration Data (12 characters, 6 bytes)

  - 4094C001C5C8

- ORP B Calibration data (12 characters, 6 bytes)

  - C0FDC0013C14

- Prior to using one of the above sensors, the calibration data must be uploaded into the Arduino using the IMport command for the sensor being used.

- Repeat this process once per year.

## DO Probe

- Pages 9-12 in the Dissolved Oxygen Manual

- For now, we are doing the single point calibration (we do not need to be calibrated below 1mg/L).  This procedure is described on page 10 of the manual.

- Connect the sensor to the EZO board and ensure connection to the arduino.

- Leave the sensor open to the air.

- Continuously request readings until the value stabilizes using the R command.

- Issue the CAL command once stabilized

- For sensor A, the readings stabilized.

- DO Calibration data for sensor A:

  - Note: The export commands failed.  Not enough power??

- For sensor B, the readings did not stabilize.

  - Note: The export commands failed.  Not enough power??

- We will contact Atles on these DO sensor issues.  In the meantime, only use DO Sensor A, as sensor B appears to have issues.

- This recalibration can be performed in the field when needed - as long as the sensor is dry.  (NOTE: DO not wipe it dry, let it air dry).

## pH probe

- This Calibration Comes from the Ph manual Pages 11-13

- Continuously Take Readings

- The sequence of testing is pH7, then pH4, thenpH10.

- For 7.00

  - Began With Removing Prob From Storage Cap

  - Rinse Probe

  - Remove the top of Ph Pouch

  - Set the Ph Into pouch until Reading Stabilize

  - Solution Is Only Valid For 20 Minutes after Discord

  - Once Reading Stabilize Send Cal,Mid,7

- Onto 4.00

- The sequence of testing is pH7, then pH4, then pH10.

- For 4.00

  - Began With Removing Probe From Storage Cap

  - Rinse Probe

  - Remove the top of Ph Pouch

  - Set the Ph Into pouch until Reading Stabilize

  - Solution Is Only Valid For 20 Minutes after Discord

  - Once Reading Stabilize Send Cal,Mid,7

## EC Probe

- Calibration Comes from page 12-16

  - Two Point Calibration (plus dry calibration)

    - Solutions are

    - 12,880us Atlas Scientific

    - 80,000us Atlas Scientific

- Do not wet the probe before Calibration

- Set Probe Type to the right voltage k 1.0 is the default

  - We queried the EZO setting and it is set to k=1.0

- Cal Dry is the first one take a cal in the air

  - No Continuous Reading to set cal,dry

  - The command failed (pending) until Mark extended the command wait time for calibration to 2 seconds.

- Calibration Results

  - Cal Probe A Low (12,880uS)

    - Calibration fluid 23 degrees Celsius - no need to compensate (within 5 degrees)

    - Waited two minutes - Issued the Cal,low,12880 command.  We queried the cal (Cal, ?), and were told it was calibrated to dry only (not a one point cal).  This is an issue. It means the calibration did not take place.

    - We tried to do the second (high) calibration.  Once settled, we issued the cal,high,80000 command.  Cal,? Responded with “?cal,0”, again meaning the calibration did not take.

    - Tried the second sensor, same result.

    - Tried the second EZO board.  Same result.

    - CRAIG TOOK THE ACTION TO CONTACT ATLAS

## Calibration results tables

Build table to collect calibration results